Modular Power Supply Controller

Setup and Calibration Notes

National Superconducting Cyclotron Laboratory, MSU

## Created:  Mark Davis, 11/12/2008      Last Updated:  Mark Davis, 11/17/2008

##### EE Document #: 211-0060-##-##_00.doc

# Introduction

The NSCL Modular Power Supply controller (ModularPS) is used in several variations requiring special care to make sure it is properly configured before proper calibration can done.  This paper describes the various configurations currently in use and outlines the proper procedure for calibrating in each case.

This paper should be used in conjunction with the following EE department docucments:

  - Automatic Controller Configuration: 111-0016-01-01_##.doc

  - Setup and Calibration: 111-019-01-01_##.doc

WARNING:  This device contains hardware-configurable parameters in addition to the standard Device ID switches.  Setting the ID switches alone is NOT sufficient to insure proper operation when swapping controllers!  Any change to the hardware components (INCLUDING altering settings on trimmer pots, etc) requires the controller to be recalibrated!

Overview of the issues:

The Modular Power Supply design allows up to 10 separately monitored supplies to be connected to a single load, allowing each supply to provide a roughly equal share of the total load current.

To aid in this flexibility, the design allows 3 different means for measuring the current through the load, and hardware-configurable gains that can be set based on the maximum current through and voltage across the load.

Proper calibration of a Modular Power Supply controller requires that these settings be taken in to account.  Any change in these settings WILL require the controller calibrations to be redone.

To make matters even more complicated, some of the variations consist of two supplies in parallel in the controller or in an external bulk-supply box, but only ONE of parallel supplies connected to the bulk supply inputs, so the controller cannot monitor half of the supplies!

Following is a description of how to properly calibrate each of the currently used configurations:

(NOTE:  The scripts used in these directions are located at:  \\icarus\ctlrdata\defaults)

10/40 (10 Volts / 40 Amps)

Description:

- A single bulk supply inside the controller

- An internal shunt for reading the load current

- An internal HOLEC for reading the bulk supply current

- Readable range for IFDBK (load current):  +/- 40 amps

- Readable range for VOUT (load voltage):  +/-15 volts

- Readable range for bulk supply current:  0 to +50 amps

- Readable range for bulk supply bias voltage:  0 to +15 volts

Calibration process:

- Apply the default settings and calibration values using the following command in a telnet session:

sh /srfs/defaults/modps_10-40_I1

- Set VR3 (the shunt gain) as follows:

  - Disconnect the leads from the shunt and apply a voltage to them equal to the voltage you would get with a current of 40 amps running through the shunt.

  - Adjust trimmer pot VR3 so the center pin of JP8 (IFDBK) reads just below 10 volts.

  - Remove the voltage source and reconnect the shunt leads.

- Test controller with a resistive load and make sure it operates correctly and that all settings and readings are reasonably close to correct.

- Use the normal calibration process described in the “Setup and Calibration” document to do a BENCH calibration (adccalib or daccalib) of the following signals:

cur5, bias5, ifdbk, vout, iset

- No field calibrations of the analog signals should be done for this configuration.

- When finished testing and calibrating, set the load type to Inductive using the LCD display or the following commands:

reg loadtype 1

save devdata

NOTE:  Each load requires some tuning of the device-specific parameters (the sinkEndAmps setting in particular).  This must be done the first time a controller is connected to a new load, but once done, the settings will be saved for the load’s Device ID and restored to any ModularPS controller configured with the same Device ID.  See the notes on field calibration for how to tune the settings for an inductive load.

10/125 (10 Volts / 125 Amps)

Description:

- A single bulk supply inside the controller

- An internal HOLEC for reading the load current

- An internal HOLEC for reading the bulk supply current

- Readable range for IFDBK (load current):  +/- 125 amps

- Readable range for VOUT (load voltage):  +/-15 volts

- Readable range for bulk supply current:  0 to +160 amps

- Readable range for bulk supply bias voltage:  0 to +15 volts

Calibration process:

- Apply the default settings and calibration values using the following command in a telnet session:

sh /srfs/defaults/modps_10-125_I1

- Test controller with a resistive load and make sure it operates correctly and that all settings and readings are reasonably close to correct.

- Use the normal calibration process described in the “Setup and Calibration” document to do a BENCH calibration (adccalib or daccalib) of the following signals:

cur5, bias5, ifdbk, vout, iset

- No field calibrations of the analog signals should be done for this configuration.

- When finished testing and calibrating, set the load type to Inductive using the LCD display or the following commands:

reg loadtype 1

save devdata

NOTE:  Each load requires some tuning of the device-specific parameters (the sinkEndAmps setting in particular).  This must be done the first time a controller is connected to a new load, but once done, the settings will be saved for the load’s Device ID and restored to any ModularPS controller configured with the same Device ID.  See the notes on field calibration for how to tune the settings for an inductive load.

10/250 (10 Volts / 250 Amps)  Standard

Description:

- 2 bulk supplies in parallel inside the controller

- An internal HOLEC for reading the load current

- An internal HOLEC for reading the bulk supply current

- Readable range for IFDBK (load current):  +/- 250 amps

- Readable range for VOUT (load voltage):  +/-15 volts

- Readable range for bulk supply current:  0 to +160 amps (bench), 0 to +320 amps (field)

- Readable range for bulk supply bias voltage:  0 to +15 volts

Calibration process:

- Apply the default settings and calibration values using the following command in a telnet session:

sh /srfs/defaults/modps_10-250_I2

- Test controller with a resistive load and make sure it operates correctly and that all settings and readings are reasonably close to correct.

- Use the normal calibration process described in the “Setup and Calibration” document to do a BENCH calibration (adccalib or daccalib) of the following signals:

cur5, bias5, ifdbk, vout, iset

- No field calibrations of the analog signals should be done for this configuration.

- When finished testing and calibrating, set the load type to Inductive using the LCD display or the following commands:

reg loadtype 1

save devdata

NOTE:  Each load requires some tuning of the device-specific parameters (the sinkEndAmps setting in particular).  This must be done the first time a controller is connected to a new load, but once done, the settings will be saved for the load’s Device ID and restored to any ModularPS controller configured with the same Device ID.  See the notes on field calibration for how to tune the settings for an inductive load.

10/250 (10 Volts / 250 Amps)  Modular, Internal HOLEC

Description:

- 1 to 10 external bulk supplies with (2) 125A supplies in parallel inside each box

- An internal HOLEC for reading the load current

- A voltage from each bulk supply box indicating the current flowing through 1 of the supplies in the box

- Readable range for IFDBK (load current):  +/- 250 amps (per bulk supply box)

- Readable range for VOUT (load voltage):  +/-15 volts

- Readable range for bulk supply current:  0 to +160 amps (bench), 0 to +320 amps (field)

- Readable range for bulk supply bias voltage:  0 to +15 volts

Calibration process:

- Apply the default settings and calibration values using the following command in a telnet session:

sh /srfs/defaults/modps_10-250_IH_XB

- Test controller with a resistive load and one bulk supply box connected to the first supply connector to make sure it operates correctly and that all settings and readings are reasonably close to correct.

- Connect the one bulk supply to each of the bulk supply connectors and test the supply each time to make sure each bulk supply input works properly.

- Using the normal calibration process described in the “Setup and Calibration” document to do a BENCH calibration of the bulk current (cur1-cur10) and bias voltage (bias1-bias10) readings (i.e. use the adccalib command).   Use an accurate current source to calibrate the bulk current readings (try using a range of +/- 50ma), and an accurate voltage source to calibrate the bias voltage readings.

- Connect the bulk supply to the first bulk supply connector and use the normal calibration process described in the “Setup and Calibration” document to do a BENCH calibration (adccalib or daccalib) of the following signals.

ifdbk, vout, iset

- Connect the bulk supply to each of the bulk supply connectors and do a FIELD calibration (readcalib) of the corresponding current inputs for each bulk current reading (cur1-cur10).

- When finished testing and calibrating, set the load type to Inductive using the LCD display or the following commands:

reg loadtype 1

save devdata

NOTE:  Each load requires some tuning of the device-specific parameters (the sinkEndAmps setting in particular).  This must be done the first time a controller is connected to a new load, but once done, the settings will be saved for the load’s Device ID and restored to any ModularPS controller configured with the same Device ID.  See the notes on field calibration for how to tune the settings for an inductive load.

20/25 (20 Volts / 25 Amps)

Description:

- A single bulk supply inside the controller

- An internal shunt for reading the load current

- An internal HOLEC for reading the bulk supply current

- Readable range for IFDBK (load current):  +/- 25 amps

- Readable range for VOUT (load voltage):  +/-24 volts

- Readable range for bulk supply current:  0 to +32 amps

- Readable range for bulk supply bias voltage:  0 to +24 volts

Calibration process:

- Apply the default settings and calibration values using the following command in a telnet session:

sh /srfs/defaults/modps_20-25_I1

- Set VR3 (the shunt gain) as follows:

  - Disconnect the leads from the shunt and apply a voltage to them equal to the voltage you would get with a current of 25 amps running through the shunt.

  - Adjust trimmer pot VR3 so the center pin of JP8 (IFDBK) reads just below 10 volts.

  - Remove the voltage source and reconnect the shunt leads.

- Test controller with a resistive load and make sure it operates correctly and that all settings and readings are reasonably close to correct.

- Use the normal calibration process described in the “Setup and Calibration” document to do a BENCH calibration (adccalib or daccalib) of the following signals:

cur5, bias5, ifdbk, vout, iset

- No field calibrations of the analog signals should be done for this configuration.

- When finished testing and calibrating, set the load type to Inductive using the LCD display or the following commands:

reg loadtype 1

save devdata

NOTE:  Each load requires some tuning of the device-specific parameters (the sinkEndAmps setting in particular).  This must be done the first time a controller is connected to a new load, but once done, the settings will be saved for the load’s Device ID and restored to any ModularPS controller configured with the same Device ID.  See the notes on field calibration for how to tune the settings for an inductive load.