Beam Current Monitor (BCM)

Software Requirements

National Superconducting Cyclotron Laboratory, MSU

## Created:  Mark Davis, 03/17/2006      Last Updated:  Mark Davis, 07/25/2006

# Introduction

The NSCL Beam Current Monitor (BCM) is a device used to monitor two separate current flows.  Its primary use is to measure intensity of the particle beam at numerous points along the beam-line.

Basic functional capabilities:

- Measure 1 or 2 independent current flows in the range of 1 milliAmp to 1 picoAmp

- A Front-End board with 8 different gain settings allowing 8 distinct ranges:

+/-1mA, +/-100uA, +/-10uA, +/-1uA, +/-100nA, +/-10nA, +/-1nA, +/-100pA

- Each hardware range supports 2 settings – one that matches the hardware, and a “software” range that is 30% of the hardware setting (this is for historical purposes)

- Sampling rate of 100 times per second (NOTE:  The delay between changes in the input and resulting changes to the reading increases rapidly as more sensitive ranges are used).

- Hardware filtering in the input amplifiers that can be enabled/disabled via software (forced on for more sensitive ranges)

- Software filtering in the form of averaging over many samples

- Remote control and monitoring via ethernet connection

- BCM unit-specific calibration/configuration values stored in the controller, automatically backed up by the control system.

Configurable parameters:

- Gain and offset for each range of each of the Front End boards.  These values are normally set only by a calibration process using an accurate, computer-controlled current source.

Values settable at runtime:

- Operating range (8 hardware ranges, each divided in to 2 “software” ranges, for a total of 16 ranges).  Each hardware range varies from adjacent ones by a factor of 10 (10 times more sensitive and 1/10 as large).

- Number of samples to use when averaging values to reduce noise (specified as an index in to a table – the default is index 0 for 10 samples).

- Turn on/off the “Front End”.  The Front End is the board connected to the input leads.  It contains amplifiers with settable gains to boost the signal used to measure the current source.  Turning “ON” the Front End connects the amplifiers to the input signal so it can be measured.

- Enable/disable hardware filter.  Each Front End board contains capacitors that can be connected or disconnected to the OpAmp amplifier feedback loop.  Connecting them provides some hardware noise filtering.

- Turn on/off test current for each channel (for devices with “solid state” switches in the Front End board).

# Required Operational Logic

- Calibration data (gain and offset) for each range is stored in flash.  This information is read from flash on startup but not overwritten without an external command.

- The state of analog and digital inputs are updated 100 times a second and can be read via the control system at any time over ethernet.

- Each channel on a BCM is in one of 2 states:

  - RTZERO_CALIB:  Busy configuring each range so it reads zero for the existing input current.  Described in more detail below.

  - RUNNING:  Updating the reading and flags for the current range setting.  Described in more detail below.

- The RTZERO_CALIB state:  This is the state a BCM channel is in when it is configuring each range to read 0 for the input current.  The process works as follows (this is done for each range):

  - Wait 500ms since the last change to the channel settings that could affect the reading (e.g. range, averaging index, filter, etc).  This gives the last change time to have a measurable effect on the reading and “variance” values (variance being a measure of how much the value has changed over the last second).

  - Continue to wait as long as the amount of variation in the reading keeps decreasing.  This is to insure that the reading has “settled” after the last change.

  - Wait an additional 2 seconds.  This insures we are not grabbing a “best-case” reading, and that the averaged value is based only recent values.

  - The difference between the unbiased (no offset applied) average voltage reading  and the bench-top calibration zero-level offset is saved as the run-time zero-level offset.  This is the value used to adjust the reading before computing a current value using the bench-top calibrated offset and gain.

Once all the ranges have been completed, the channel is set to the RUNNING state.

- The RUNNING state:  This is the state a BCM channel is most often in.  In this state it accepts all commands and keeps track of where the reading is relative to the limits of the active range.  Based on the voltage level of the analog input from which the reading is computed, the running state will report one of 3 conditions for the reading (each of these percentages applies in both the positive and negative direction):

  - BCM_OVERRANGE_OK:  The voltage reading is within 92% of the maximum readable voltage.

  - BCM_OVERRANGE_READY:  The voltage reading is between 92% and 99% of the maximum readable voltage.

  - BCM_OVERRANGE_ERROR:  The voltage reading is greater then 99% of the maximum readable voltage.

Note that these limits are NOT based on the nominal current reading limits for each range, but what the hardware can actually read reliably.  What this means is that the current readings that correspond to the hardware limits will vary based on how big an offset there is from the logical 0-level reading (see the RTZERO_CALIB state for a description of how the offsets are determined).

Example:  A voltage range of +/-10V and an offset of +5V (meaning at 5V, none of the current through the input is part of the signal we are interested in, so it should read 0) means the larget positive current the range can read is 50% of the nominal limit, but the largest negative current is 150% of the nominal limit.

- On startup, each BCM channel is initialized as follows:

  - The calibration data is read from flash memory and a “sanity” check is done on the values for each range.  The run-time (i.e. in RAM) values will be set to defaults for any range that fails the sanity check and the range flagged as “BAD” (BAD meaning the bench-top calibration values in flash appear to be bad).
NOTE:  The bench-top calibrated gain and offset values stored in flash are NEVER altered without a specific command to do so (normally sent by the bench-top calibration application).

  - Clear OVER-RANGE status bit

  - Turn on Front-End (i.e. connect input amplifiers so current flow can be read)

  - Disable the hardware filter

  - Turn off the Test Current (for boards that support software control of this feature)

  - Set range to 0 (1mA max, the least sensitive range)

  - Set # samples to average over to default setting (index 0, 10 samples)

  - Enter the RUNNING state

- Each BCM channel supports the following commands:

NOTE:  All but the AbortCalib command is ignored ruing a run-time 0-level calibration.

  - Init:  Perform the startup sequence for the channel.

  - DoCalibrate:  Sets the channel to the RTZERO_CALIB state as described above.

  - AbortCalib:  Abort current calibration process.

  - FECalibPoint:  Save the specified current and corresponding voltage reading to use in computing the offset and gain for the current range.

  - RangeUp: Change the range to the next most sensitive one

  - RangeDown:  Change the range to the next least sensitive one

  - AvgIdxUp:  Increase the index in to the table of # of samples to average over

  - AvgIdxDown:  Decrease the index in to the table of # of samples to average over

  - FrontEndOn:  Connect the Front-End amplifier board to its input so that the current going through it can be measured.  NOTE that the FrontEnd should normally be OFF when the test current is enabled.

  - FrontEndOff:  Disconnect the Front-End amplifier board from its input (disables the ability to read the current going through the input).  NOTE that the FrontEnd should normally be OFF when the test current is enabled.

  - FilterOn:  Turn on the filter in the Front-End’s amplifier circuit.

  - FilterOff:  Turn off the filter in the Front-End’s amplifier circuit.

  - ClearOffsets:  Clears all the run-time zero-level offsets (sets them all to 0).

  - TestCurOn:  Connect test current (150nA) to input (for models that support this).

  - TestCurOff:  Disconnect test current from input (for models that support this).

  - SaveCalib:  Save the Front-End calibration data to flash (normally used only by the application that manages the bench-top calibration process).

- The ACTIVITY LED is on the outside of the box, and is used to check the overall state of the controller.  Normally the LED blinks at approximately 8 Hz.  When calibrating, the LED blinks at approximately 4 Hz.  If there is a hardware problem, it blinks at approximately 1Hz.

# Optional Operational Logic

Auto-Ranging:  Older models had a version of this feature, but the implementation was fairly simple and did not meet expectations and so was disabled in later models.  Also, the operators often monitor the reading via a meter, and having the meter jump around due to the range automatically changing can make beam tuning difficult.  If re-implemented, this feature must support being enabled and disabled.

The feature is meant to allow the BCM to change ranges on its own to avoid the following conditions:

  - The voltage is close to or at the maximum readable value for the active range.  When close to the maximum, readings can be clipped, loosing some of the data.  When at the maximum, all information is lost since the true value is never within the range that the hardware can manage.

  - The voltage is close to 0.  In this case, the signal to noise ratio is poor and the true values can be lost in the noise, greatly reducing the accuracy of the resulting data.

The Auto-Ranging feature would track these conditions and, based on algorithms to decide when it is appropriate, automatically increase or decrease the active range to avoid such conditions.

To avoid having the range change too often or produce a “jump” in the reading when the range is changed, a number of techniques might be used.  Here are 2 likely possibilities:

- Use a weighted “accumulated time” value to decide when to change ranges.  This would be computed as a continuous function where values close to 0 or the maximum increase the total rapidly, values in the middle of the range decrease it rapidly (although never below 0), and values in between the middle and 0 or maximum do not change it at all.

This would avoid the likelihood of numerous changes for signals that vary a lot near a fixed dead-band range and/or have occasional jumps in amplitude.  Essentially it provides a “dampened” response in terms of deciding to change ranges based on changes in the signal.  NOTE:  The same or similar behavior can be probably be achieved by simply keeping a running average of the signal and applying a simple dead-band to the averaged value.  Adjusting the # of samples in the average would control how quickly it responds to changes.

- Temporarily increase the # of values that are used to compute an averaged reading whenever the range is automatically changed.  This would help smooth out sudden jumps in the reading because of the difference in absolute accuracy between the previous range and the new one.  Like the dampened approach to choosing when to change the range, the # of values in the average could also be varied continuously as a function of how much time has passed since the switch was made.