Modular Power Supply Controller

Software Requirements

National Superconducting Cyclotron Laboratory, MSU

## Created:  Mark Davis, 05/22/2007      Last Updated:  Mark Davis, 09/25/2007

##### EE Document #: 211-0060-02-01_00.doc

# Introduction

The NSCL Modular Power Supply controller (ModularPS) is a rack-mounted module used to monitor and control a 4-quadrant power supply comprised of up to 10 inter-dependent analog power supplies.

Basic functional capabilities:

- Control and monitor up to 10 power supplies and their combined output

- Monitor the state of each supply independently and enable/disable any combination of them

- Display state information and accept commands from remote clients and a front panel interface consisting of an LCD text/graphics display, a knob, and push-buttons

- Remote control and monitoring via Ethernet connection

- Controller and Device-specific configuration and calibration information stored in a small file system on the controller, with backup copies kept on a remote server.  A unique ID is assigned to the device (or group of devices) managed by a controller, allowing a new controller to be swapped for an existing one and configure itself from the backup copy of the Device-specific information stored by the older controller.  For details, see the latest version of the following EE department documents:

  - Automatic Controller Configuration: 111-0016-01-01_##.doc

  - Setup and Calibration: 111-019-01-01_##.doc

WARNING:  This device contains hardware-configurable parameters in addition to the standard Device ID switches.  Setting the ID switches alone is NOT sufficient to insure proper operation when swapping controllers!  Any change to the hardware components (INCLUDING altering settings on trimmer pots, etc) requires the controller to be recalibrated!

Functional overview:

The Modular Power Supply is a controller that monitors and controls up to 10 interdependent power supplies all connected to a single load (normally an inductive load, a large electromagnet).   The output of the power supplies are connected to a “bridge” consisting of FET’s on the “top” half of the bridge, BJT’s on the “bottom” half of the bridge, with the load connected in the middle.   The FET’s and BJT’s are enabled in pairs to control the polarity of the load.

The drivers for the FET’s are under direct software control, while the drivers for the BJT’s are controlled by the hardware, naturally following the state of the load.  The output of the feedback control loop circuit controls which BJT driver is active.  Reading this signal allows the software to enable the correct polarity FET driver.

The +/- polarity designations given to the FET and BJT driver outputs is such that the +FET and +BJT may be enabled together to drive the load in one polarity, and the –FET and –BJT enabled together to drive the load in the opposite polarity.

To prevent shorting the output of the supplies through the bridge, only one FET driver should be active at a time, and the polarity of an active FET driver must match the polarity of the currently active BJT driver.

Configurable Engineering parameters:

These are values that are normally configured when a controller is installed and tested with a new set of power supplies.  The software-configurable values are included in the Device-specific values stored in a file on the controller (with backup copies on a remote server).

Modification of the following parameters (both hardware and software) should be limited to engineering staff with knowledge on how to properly configure them.

Software Settings:

- Gain and offset for each of the analog inputs and outputs:  A table of values that determines the conversion between engineering units and analog input and output values.  Set using firmware calibration commands while the equipment is in operation.

- vlimPlus:  The maximum positive voltage across the load

- vlimMinus:  The maximum negative voltage across the load

- rampRate:  The rate at which the active setpoint changes until it reaches the requested setpoint (in Amps/second)

- vctrlHyst:  The +/- hysteresis threshold that VCTRL must cross to change the perceived polarity of the BJT drivers (in volts)

- maxErrTime:  The maximum contiguous time normal operating parameters are allowed to be exceeded before an Interlock condition is asserted (in ms)

- maxVctrl:  The maximum value VCTRL is allowed to reach (for a period exceeding maxErrTime) before an Interlock condition is asserted.  This is provided as a safeguard during testing of software changes that might damage the FET/BJT bridge in the power supplies.  For normal operation it must be set high enough to avoid causing a false  interlock condition (in volts)

- forceBJTtime:  The amount of time that the polarity of the active BJT is ignored after SINK mode ends.  This stops the control loop from causing oscillations due to the brief spike in VCTRL that happens when the FET's are turned back on at the end of a SINK interval (in ms)

- maxCurDiff:  Currently ignored.  Will be used in the future to check for problems with the individual bulk supplies.

- maxVoltDiff:  Currently ignored.  Will be used in the future to check for problems with the individual bulk supplies.

- maxDiffSec:  Currently ignored.  Will be used in the future to check for problems with the individual bulk supplies.

- knobSelectRate:  Controls how much the knob on the front panel has be turned to change what value is selected (in ticks)

- maxAmpSet:  The maximum allowable setpoint (in amps)

- minPSonTime:  The minimum amount of time the bulk supplies must be on before normal operation is allowed (in ms)

- minFEToffTime:  Currently ignored.  Possibly used again in the future to force a delay between turning FET's off and on (in ms)

- minBJTstableTime:  Minimum amount of time the polarity of the active BJT must not change before the last known stable state of the BJT's is updated (in ms)

- rampThreshold:  The minimum difference between ISET and IFDBK values that indicates the supply is RAMPING or SINKING.  NOTE:  This value is in volts, NOT amps!!!  This avoids a chicken-and-egg problem of how to perform field-level calibrations on a controller that requires field-level calibrations to operate.

Hardware Settings:

- Device ID switches:  Must be set to the ID assigned to the device(s).   Used by the control system to locate and configure controllers, and by controllers when saving or restoring configuration data.

- IFDBK jumper (JP1):  Enables/Disables the IFDBK portion of the hardware feedback loop.  Removing this jumper requires the software control loop to manage the ISET value based not just on the active set-point, but upon the value of IFDBK as well.  This mode of operation is not included in the firmware at this time.

- ISELECT jumper (JP2): Sets the source of the IFDBK signal.  The options are:

  - SH+/SH-:  Amplified value from across a shunt

  - EXT. HOLEC:  Un-amplified signal from an external current transformer

  - Internal HOLEC:  Amplified signal from an internal current transformer

- Gain on SH+/SH- input:  Set by a trimmer pot (VR3)

- Gain on HOLEC input.  Set by jumpers for 6 resistors (JP2-JP7).

- Gain and frequency response for internal VCTRL signal.  Set by 2 resistors and a capacitor (R67, R68, C68).

- Offset for BJT voltage.  Set by a trimmer pot (VR1) (may be removed in future version).

Inputs and Outputs, Operating Parameters

# Software Inputs/Settings

- On/Off

- Current set-point

- Reset Interlock

# Software Outputs/Readings

- Current value of all software inputs/settings

- Current operating state

- Load voltage and current

- HOLEC state (OK or BAD)

- External Interlock

- State of FET and BJT drivers

- State of each supply (Bias Voltage, Current, On/Off, connected, temp OK)

### Hardware Inputs

- Analog readings for each bulk supply:

  - HECT (Hall Effect Current Transducer): Current total for a single supply

  - BSV (Bulk Supply Voltage):  Bias voltage for a single supply

- Digital inputs for each bulk supply:

  - ON: On/Off state for a single supply

  - TMP (Temperature): Temperature OK

  - RC (Roll Call):  Presence/Absence of a supply

- Analog readings for combined operation of all bulk supplies:

  - IFDBK: proportional to current through the load (from shunt or HOLEC)

  - VOUT: equal to the voltage across the load

  - VCTRL: Control Voltage: Output of hardware feedback control loop.  Controls BJT drive and polarity

  - VLIM: Analog signal activated when VOUT exceeds VLIM_POS or VLIM_NEG (see Hardware Outputs).  Used in hardware feedback loop to limit voltage across the load.

- Digital Inputs not related to the individual bulk supplies:

  - HOLEC_OK:  Indicates that a HOLEC current transformer is connected

  - EXT_INTLK: External Interlock

- Front panel Buttons:

  - Reset (controller):  Causes embedded controller to reboot

  - Emergency Off:  Hardware override; Forces Off condition for all supplies

  - Start, Stop:  Normal On/Off control

  - Reset (interlock):  Attempts to clear the Interlock condition

  - Select, Set  (for manipulating elements on LCD display)

- Front panel Knob (for selecting and modifying values shown on LCD)

#### Hardware Outputs

- Digital outputs for each supply:

  - ON: On/Off control for a single supply

- Analog outputs for controlling combined effect of all supplies:

  - ISET: Current Set-point

  - VLIM_POS: Sets Positive Limit for voltage feedback (inverted and scaled)

  - VLIM_NEG: Sets Negative Limit for voltage feedback (inverted and scaled)

- Digital Outputs for combined effect of all supplies

  - FET_NEG:  Enable Negative FET driver

  - FET_POS:  Enable Positive FET driver

  - NEG_LED:  Indicate Negative FET driver is enabled

  - POS_LED:  Indicate Positive FET driver is enabled

  - ACTIVITY:  Blink LED at different rates for normal operation or problems

- Misc front panel connections:

  - LED’s:  Activity, (Optional: FET+, FET-)

  - Text/graphics LCD display

  - Serial TTL Diagnostic/Programming/Console port

# Theory of Operation

NOTE: In the following discussion, references to the “set-point” generally refer to the “active” set-point (the one currently driving the hardware), NOT the value toward which the active set-point is being ramped.

- Signals used by the software and/or hardware control-loops:

  - ISET:  Set by software, it provides the set-point for the feedback control circuit.

  - IFDBK:  A voltage proportional to the current flowing through the output circuit.  It is one of the inputs to the feedback control circuit and is read by the software.

  - VLIM_POS: Set by software, this limits the positive voltage across the load.

  - VLIM_NEG: Set by software, this limits the negative voltage across the load.

  - VOUT:  The voltage across the load.  Read by the software.

  - VLIM:  Analog signal activated when VOUT exceeds VLIM_POS or VLIM_NEG.  Read by the software and used in the hardware feedback loop to limit voltage across the load.

  - FET_POS and FET_NEG:  Software controlled outputs for opening or closing the FET-controlled circuit pathways.

  - VCTRL:  The output of the feedback control circuit.  It controls which BJT driver is active and regulates the current flow.

- Basic operation of the software control loop (see “Operational Logic” below for details):

The ModularPS controller supports up to 10 "bulk" power supplies.  The output is normally connected to an inductive load (a magnet) through a bridge made up of BJT's and FET's.  The BJT's are used to set the output level of the supply, while the FET's turn the bridge on or off.

The controller has an analog feedback loop that uses the following signals:

ISET:  Active set-point for the analog feedback loop

IFDBK: The current through the load

VOUT:  The voltage across the load

VLIM_NEG:  The largest negative voltage allowed across the load

VLIM_POS:  The largest positive voltage allowed across the load

VLIM:  Analog signal activated when VOUT exceeds VLIM_POS or VLIM_NEG

The combination of ISET and IFDBK provides normal regulation of the load current, while the VLIM value imposes a limit based on the voltage across the load.  The combined result (VCTRL) is used to bias the BJT's which control the combined power output of the bulk supplies.

The software's role in the control loop is to monitor the state of the load and the BJT's and determine when one side of the bridge should be turned on via the positive or negative FET driver.  The VCTRL, IFDBK, and ISET values provide the needed information as follows:

VCTRL < 0 means the Positive BJT’s are enabled.  VCTRL > 0 means the Negative BJT’s are enabled.

Define curChg = abs(IFDBK[t])) – abs(IFDBK[t-1]), where IFDBK[t] and

IFDBK[t-1] are 2 recent values for IFDBK (t-1 being the earlier one).

If abs(curChg) is below some threshold, then the load is in REGulation mode.

If curChg > 0 and exceeds some threshold, then the load is in RAMP mode.

If curChg < 0 and exceeds some threshold, then the load is in SINK mode.

Define ctlErr = ISET - IFDBKB.

If abs(ctlErr) is below some threshold, then the supply is in REGulation mode.

If abs(ctlErr) exceeds some threshold, and the polarity of the active BJT driver is equal to the polarity of ctlErr, then the supply is in RAMP mode.

If abs(ctlErr) exceeds some threshold, and the polarity of the active BJT driver is NOT equal to the polarity of ctlErr, then the supply is in SINK mode.

When the supply is in RAMP or REG mode, the bridge should be turned on and the polarity of the enabled FET driver normally must match the active BJT driver.

When the supply is in SINK mode, the FET's are disabled to avoid dumping lots of power through the bridge.

Note that it is possible for the load to be in SINK mode while the supply is in RAMP mode (in which case the supply is actually sinking power from the load).

When the supply mode changes from SINK back to RAMP or REG, the appropriate FET must be re-enabled.  However, this is where a bit of a problem occurs.

When the FET’s are re-enabled, VCTRL spikes, causing the polarity of the active BJT to change back to their SINK mode state.  If the software attempts to keep the active FET in sync with the active BJT at this time, the result is oscillation.

To avoid this problem, the changes in the BJT state must be ignored for a (configurable) period time after coming out of SINK mode.  This avoids having the control loop oscillate between SINK and RAMP or REG mode.  Only a change in polarity of the ctlErr value (ISET - IFDBKB) will shorten this period and cause the software to resume using the BJT state to recognize when the supply enters SINK mode again.

- Normal operating conditions:

  - All bulk supplies are ON.

  - The current (HECT) and voltage (BSV) readings for all supplies are within the (configurable) tolerance values.

# Operational Logic

- Real-Time main application process that runs at the highest priority:

  - Handles all controller-specific initialization and then runs a continuous loop.

  - The loop is executed every 10ms and be runs from start to finish with minimal interruptions

  - At the end of the loop, the process blocks until the start of the next 10ms interval, at which time it is resumed with minimal latency to execute the loop and block again.

  - The loop logic normally does 3 things:

    - Read the current state of all hardware input signals

    - Update the state of all hardware output signals

    - Update the state of the controller

  - Updating the state of the controller takes in to account the following:

    - The current state of the hardware inputs and outputs

    - The current logic state of the controller

    - Requests to change the state of the controller (implied by the current state of the hardware, or received from an external client via a communications channel).

  - Any changes to the desired state of the hardwired output signals are applied at the beginning of the next interval, just after reading the inputs.

- The controller supports the following logic states:

  - INIT:  The application process has not finished initializing yet.

  - OFF:  All supplies disabled

  - ON:  The bulk supplies are ON and the FET drivers are turned on and off as required for normal operation.

  - INTERLOCK:  An interlock condition (external or internal) has occurred.  Same as the OFF state, with the addition of the Interlock status, which must be reset before an ON command will be accepted.

- Processing common to all states:

  - Count the # of supplies that are connected

  - Update values as needed based on last-read state of inputs

  - Process event messages

- INIT state:

  - Initialize all outputs to be consistent with the OFF state (see below)

  - Do processing common to all states

- OFF state:

  - Turn off all the bulk power supplies

  - Disable both FETs

  - Reset values used in the ON state logic:

psReady = false;  psOnTime = lastBJTChgTime = 0;

  - Do processing common to all states

- ON state:

NOTE:  For polarity flags, TRUE means Positive and FALSE means Negative.

  - Determine if the power supplies have been on long enough to be in a stable state:

if (timeSince(psOnTime) > minPSonTime)  psReady = true;

  - Determine the polarity and magnitude of the error in the load current:

lastCtlErr = ctlErr;  ctlErr = ISET - IFDBK;

polErr = (ctlErr >= 0.0);  absErr = FABS(ctlErr);

  - Determine polarity of BJT drivers (VCTRL < 0 means Positive BJT’s are enabled):

polBJT = (VCTRL <= 0);   // using no hysteresis

if (VCTRL < -vctrlHyst)  hystPolBJT = TRUE;    // pos polarity using hysteresis

if (VCTRL >  vctrlHyst)  hystPolBJT =  FALSE;  // neg polarity using hysteresis

  - Determine if the load is in SINK mode.  To avoid oscillation, a delay is inserted between SINK periods, which can be reset by a change in the polarity of the control error (indicating the load needs to be driven in the opposite polarity).

ramping = (absErr > rampThreshold);  // is difference well beyond noise level?

if (SIGN(ctlErr) != SIGN(lastCtlErr))  {  // new dir to drive the load in?

sinkMode = false;  sinkModeEndTime = 0;  }  // reset SINK mode flags

else  {

sinkMode |= ((polBJT != polErr) AND ramping AND !sinkModeEndTime);

if (sinkMode AND (polBJT == polErr))  {

sinkMode = false;  tempPolBJT = polBJT;  sinkModeEndTime = GetMS(); }

}

  - Set the state of the FET drivers.  If turning them on due to the end of a SINK period, temporarily ignore the current state of the BJT’s (which can oscillate for a bit).

oscillationPeriod = false;

if (sinkModeEndTime)  {  // SINK mode recently ended ?

if (msSince(sinkModeEndTime) < forceBJTtime)

oscillationPeriod = true;

else

sinkModeEndTime = 0;  // allow sink mode to be reasserted in future

}

fetState = FET_OFF;

if (psReady AND !sinkMode)  {

if (oscillationPeriod) setPolFET = tempPolBJT; else setPolFET = hystPolBJT;

if (setPolFET)  fetState = FET_POS;  else  fetState = FET_NEG;

}

  - Combine output state logic values in to a single index:

if (state != STATE_ON)  outMode = OFF;  else

if (!psReady)   outMode = WAIT;  else

if (sinkMode)  outMode = (polBJT ? SINK_P : SINK_N);  else

if (ramping)    outMode = (polBJT ? RAMP_P : RAMP_N);  else

outMode = (polBJT ? REG_P : REG_N);

  - Adjust ISET based on ramp rate, current ISET value and user-supplied set-point

  - For each bulk supply, determine if the current (HECT) or voltage (BSV) values are within the configured thresholds for normal operation.  If both are within the normal ranges, then update the last known time of normal operation for the supply.

  - Disable any bulk supply whose values have consistently been outside the normal ranges for longer than the configured time limit and log an error.

  - Do processing common to all states

- INTERLOCK state:

  - Interlock status = True

  - All other values set the same as OFF state

  - Controller remains in the Interlock state until acknowledged by a Reset command at a time that no interlock condition exists.

  - The External Interlock condition is asserted any time there is a break in the external interlock loop circuit.

  - An Emergency Stop button on the front panel will be part of this circuit, and pressing the button will interrupt the circuit, asserting an interlock condition.

  - The state of the External Interlock condition will be available to the controller via one of its digital inputs.

    - When the controller sees the External Interlock condition asserted, it disables all the power supplies and leaves them off until the interlock condition is cleared and a command is received to turn the supplies back on.

    - The circuit design must NOT allow the controller to enable any of the power supplies while an External Interlock condition is in effect.

- Requests to change between states:

  - CLEAR command:  Request to change from INTERLOCK to OFF state.

    - If the controller is not in the INTERLOCK state, the state is not changed.

    - If the controller is in the INTERLOCK state, and no interlock conditions exist, it will change to the OFF state.

  - ON command:  Request to change from OFF to ON state.

    - If the controller is in the OFF state, it will change to the ON state.

    - If the controller is in the INTERLOCK state, the state will not change.

  - OFF command:  Request to change from ON to OFF state.

    - If the controller is in the ON state, it will change to the OFF state.

    - If the controller is in the INTERLOCK state, the state will not change.

- Start button:  Press by itself to assert an ON command.

- Stop button:  Press by itself to assert an OFF command.

- Reset button:  Press by itself to assert a CLEAR command.

- Emergency Stop button:  Large red pushbutton on the front panel, stable in the open or closed state.

  - The button forms part of the External Interlock circuit.  Pressing it introduces a break in the Interlock circuit, forcing an interlock and disabling all the power supplies.

  - Pulling the button outward removes the break and allows the INTERLOCK condition to be cleared via a subsequent CLEAR command.

  - All INTERLOCK state logic applies.

- Software vs Hardware regulation/feedback loop operation:  At this time, the controller does not support regulation via software.  A later version of the software may support this to allow greater flexibility in the control loop.

- Accept, process, and respond to commands.  Commands can be generated in response to the current state or changes in state of the hardware inputs and outputs, or from local or remote processes by via standard communications interfaces.

- The ACTIVITY LED is on the outside of the box, and is used to check the overall state of the controller.  Normally the LED blinks at approximately 8 Hz.  During any kind of run-time calibration, the LED blinks at approximately 4 Hz.  If there is a hardware problem, it blinks at approximately 1Hz.