Tester:   Mark Davis

SerialNo: 8

Comments: N1 Vault BMC2002

Calibrating BCM channel

Determining Zero Offsets...

Calibrating range 0

Set Point 1: -700.0E-06

Set Point 2: 700.0E-06

Gain: 132.4E-06 Amps/Volt   Offset: 0.007964 Volts

Calibrating range 1

Set Point 1: -210.0E-06

Set Point 2: 210.0E-06

Gain: 132.5E-06 Amps/Volt   Offset: 0.007964 Volts

Calibrating range 2

Set Point 1: -70.00E-06

Set Point 2: 70.00E-06

Gain: 11.59E-06 Amps/Volt   Offset: 0.007964 Volts

Calibrating range 3

Set Point 1: -21.00E-06

Set Point 2: 21.00E-06

Gain: 11.58E-06 Amps/Volt   Offset: 0.007964 Volts

Calibrating range 4

Set Point 1: -7.000E-06

Set Point 2: 7.000E-06

Gain: 1.370E-06 Amps/Volt   Offset: 0.012924 Volts

Calibrating range 5

Set Point 1: -2.100E-06

Set Point 2: 2.100E-06

Gain: 1.370E-06 Amps/Volt   Offset: 0.012924 Volts

Calibrating range 6

Set Point 1: -700.0E-09

Set Point 2: 700.0E-09

Gain: 125.9E-09 Amps/Volt   Offset: 0.022845 Volts

Calibrating range 7

Set Point 1: -210.0E-09

Set Point 2: 210.0E-09

Gain: 125.8E-09 Amps/Volt   Offset: 0.022845 Volts

Calibrating range 8

Set Point 1: -70.00E-09

Set Point 2: 70.00E-09

Gain: 12.61E-09 Amps/Volt   Offset: 0.042685 Volts

Calibrating range 9

Set Point 1: -21.00E-09

Set Point 2: 21.00E-09

Gain: 12.62E-09 Amps/Volt   Offset: 0.042685 Volts

Calibrating range 10

Set Point 1: -7.000E-09

Set Point 2: 7.000E-09

Gain: 1.155E-09 Amps/Volt   Offset: 0.012924 Volts

Calibrating range 11

Set Point 1: -2.100E-09

Set Point 2: 2.100E-09

Gain: 1.157E-09 Amps/Volt   Offset: 0.012924 Volts

Calibrating range 12

Set Point 1: -700.0E-12

Set Point 2: 700.0E-12

Gain: 137.4E-12 Amps/Volt   Offset: 0.102207 Volts

Calibrating range 13

Set Point 1: -210.0E-12

Set Point 2: 210.0E-12

Gain: 136.7E-12 Amps/Volt   Offset: 0.102207 Volts

Calibrating range 14

Set Point 1: -70.00E-12

Set Point 2: 70.00E-12

Gain: 12.70E-12 Amps/Volt   Offset: 1.079360 Volts

Calibrating range 15

Set Point 1: -21.00E-12

Set Point 2: 21.00E-12

Gain: 12.80E-12 Amps/Volt   Offset: 1.079360 Volts

Calibration of channel 0 Complete

Test started:  06/21/2002 14:03

BCM ROM Date:  2002-06-06

BCM Chan:      0

Front End:     ON

Filter:        ON

Range  0   FullScale:    1.000E-03   Offset perc/FS:  0.0%

%FS: -90.0   I: -900.0E-06A   Avg: -900.7E-06   Err:    0.08%

%FS: -50.0   I: -500.0E-06A   Avg: -500.4E-06   Err:    0.08%

%FS: -10.0   I: -100.0E-06A   Avg: -100.0E-06   Err:    0.01%

%FS:  10.0   I:  100.0E-06A   Avg:  99.78E-06   Err:   -0.22%

%FS:  50.0   I:  500.0E-06A   Avg:  499.6E-06   Err:   -0.08%

%FS:  90.0   I:  900.0E-06A   Avg:  899.6E-06   Err:   -0.04%

Range  2   FullScale:    100.0E-06   Offset perc/FS:  0.0%

%FS: -90.0   I: -90.00E-06A   Avg: -90.09E-06   Err:    0.10%

%FS: -50.0   I: -50.00E-06A   Avg: -50.06E-06   Err:    0.11%

%FS: -10.0   I: -10.00E-06A   Avg: -10.00E-06   Err:    0.04%

%FS:  10.0   I:  10.00E-06A   Avg:  9.981E-06   Err:   -0.19%

%FS:  50.0   I:  50.00E-06A   Avg:  49.96E-06   Err:   -0.07%

%FS:  90.0   I:  90.00E-06A   Avg:  89.99E-06   Err:   -0.01%

Range  4   FullScale:    10.00E-06   Offset perc/FS:  0.1%

%FS: -90.0   I: -9.000E-06A   Avg: -9.014E-06   Err:    0.16%

%FS: -50.0   I: -5.000E-06A   Avg: -5.009E-06   Err:    0.17%

%FS: -10.0   I: -1.000E-06A   Avg: -1.006E-06   Err:    0.58%

%FS:  10.0   I:  1.000E-06A   Avg:  991.9E-09   Err:   -0.81%

%FS:  50.0   I:  5.000E-06A   Avg:  4.991E-06   Err:   -0.17%

%FS:  90.0   I:  9.000E-06A   Avg:  8.997E-06   Err:   -0.04%

Range  6   FullScale:    1.000E-06   Offset perc/FS:  0.2%

%FS: -90.0   I: -900.0E-09A   Avg: -899.7E-09   Err:   -0.03%

%FS: -50.0   I: -500.0E-09A   Avg: -499.6E-09   Err:   -0.09%

%FS: -10.0   I: -100.0E-09A   Avg: -99.26E-09   Err:   -0.74%

%FS:  10.0   I:  100.0E-09A   Avg:  100.5E-09   Err:    0.47%

%FS:  50.0   I:  500.0E-09A   Avg:  500.4E-09   Err:    0.08%

%FS:  90.0   I:  900.0E-09A   Avg:  900.4E-09   Err:    0.04%

Range  8   FullScale:    100.0E-09   Offset perc/FS:  0.4%

%FS: -90.0   I: -90.00E-09A   Avg: -90.48E-09   Err:    0.53%

%FS: -50.0   I: -50.00E-09A   Avg: -50.46E-09   Err:    0.93%

%FS: -10.0   I: -10.00E-09A   Avg: -10.46E-09   Err:    4.58%

%FS:  10.0   I:  10.00E-09A   Avg:  9.545E-09   Err:   -4.55%

%FS:  50.0   I:  50.00E-09A   Avg:  49.53E-09   Err:   -0.93%

%FS:  90.0   I:  90.00E-09A   Avg:  89.56E-09   Err:   -0.49%

Range 10   FullScale:    10.00E-09   Offset perc/FS:  0.1%

%FS: -90.0   I: -9.000E-09A   Avg: -9.002E-09   Err:    0.02%

%FS: -50.0   I: -5.000E-09A   Avg: -4.995E-09   Err:   -0.10%

%FS: -10.0   I: -1.000E-09A   Avg: -998.9E-12   Err:   -0.11%

%FS:  10.0   I:  1.000E-09A   Avg:  1.005E-09   Err:    0.49%

%FS:  50.0   I:  5.000E-09A   Avg:  5.009E-09   Err:    0.19%

%FS:  90.0   I:  9.000E-09A   Avg:  9.006E-09   Err:    0.06%

Range 12   FullScale:    1.000E-09   Offset perc/FS:  1.1%

%FS: -90.0   I: -900.0E-12A   Avg: -900.4E-12   Err:    0.04%

%FS: -50.0   I: -500.0E-12A   Avg: -499.9E-12   Err:   -0.02%

%FS: -10.0   I: -100.0E-12A   Avg: -99.38E-12   Err:   -0.62%

%FS:  10.0   I:  100.0E-12A   Avg:  101.6E-12   Err:    1.59%

%FS:  50.0   I:  500.0E-12A   Avg:  501.0E-12   Err:    0.21%

%FS:  90.0   I:  900.0E-12A   Avg:  902.3E-12   Err:    0.26%

Range 14   FullScale:    100.0E-12   Offset perc/FS: 11.9%

%FS: -90.0   I: -90.00E-12A   Avg: -90.43E-12   Err:    0.48%

%FS: -50.0   I: -50.00E-12A   Avg: -50.38E-12   Err:    0.76%

%FS: -10.0   I: -10.00E-12A   Avg: -11.05E-12   Err:   10.52%

%FS:  10.0   I:  10.00E-12A   Avg:  8.962E-12   Err:  -10.38%

%FS:  50.0   I:  50.00E-12A   Avg:  50.32E-12   Err:    0.63%

%FS:  90.0   I:  90.00E-12A   Avg:  90.37E-12   Err:    0.41%

Test stopped

Test time: 8 min 15 sec

Calibrating BCM channel

Determining Zero Offsets...

Calibrating range 0

Set Point 1: -700.0E-06

Set Point 2: 700.0E-06

Gain: 132.3E-06 Amps/Volt   Offset: 0.013362 Volts

Calibrating range 1

Set Point 1: -210.0E-06

Set Point 2: 210.0E-06

Gain: 132.5E-06 Amps/Volt   Offset: 0.013362 Volts

Calibrating range 2

Set Point 1: -70.00E-06

Set Point 2: 70.00E-06

Gain: 11.60E-06 Amps/Volt   Offset: 0.008371 Volts

Calibrating range 3

Set Point 1: -21.00E-06

Set Point 2: 21.00E-06

Gain: 11.58E-06 Amps/Volt   Offset: 0.008371 Volts

Calibrating range 4

Set Point 1: -7.000E-06

Set Point 2: 7.000E-06

Gain: 1.370E-06 Amps/Volt   Offset: -0.001609 Volts

Calibrating range 5

Set Point 1: -2.100E-06

Set Point 2: 2.100E-06

Gain: 1.371E-06 Amps/Volt   Offset: -0.001609 Volts

Calibrating range 6

Set Point 1: -700.0E-09

Set Point 2: 700.0E-09

Gain: 126.1E-09 Amps/Volt   Offset: -0.166292 Volts

Calibrating range 7

Set Point 1: -210.0E-09

Set Point 2: 210.0E-09

Gain: 126.2E-09 Amps/Volt   Offset: -0.166292 Volts

Calibrating range 8

Set Point 1: -70.00E-09

Set Point 2: 70.00E-09

Gain: 12.61E-09 Amps/Volt   Offset: 0.008371 Volts

Calibrating range 9

Set Point 1: -21.00E-09

Set Point 2: 21.00E-09

Gain: 12.62E-09 Amps/Volt   Offset: 0.008371 Volts

Calibrating range 10

Set Point 1: -7.000E-09

Set Point 2: 7.000E-09

Gain: 1.156E-09 Amps/Volt   Offset: -0.011590 Volts

Calibrating range 11

Set Point 1: -2.100E-09

Set Point 2: 2.100E-09

Gain: 1.158E-09 Amps/Volt   Offset: -0.011590 Volts

Calibrating range 12

Set Point 1: -700.0E-12

Set Point 2: 700.0E-12

Gain: 137.1E-12 Amps/Volt   Offset: -0.241148 Volts

Calibrating range 13

Set Point 1: -210.0E-12

Set Point 2: 210.0E-12

Gain: 137.0E-12 Amps/Volt   Offset: -0.241148 Volts

Calibrating range 14

Set Point 1: -70.00E-12

Set Point 2: 70.00E-12

Gain: 12.60E-12 Amps/Volt   Offset: -2.761300 Volts

Calibrating range 15

Set Point 1: -21.00E-12

Set Point 2: 21.00E-12

Gain: 12.60E-12 Amps/Volt   Offset: -2.761300 Volts

Calibration of channel 1 Complete

Test started:  06/21/2002 14:19

BCM ROM Date:  2002-06-06

BCM Chan:      1

Front End:     ON

Filter:        ON

Range  0   FullScale:    1.000E-03   Offset perc/FS:  0.1%

%FS: -90.0   I: -900.0E-06A   Avg: -900.3E-06   Err:    0.04%

%FS: -50.0   I: -500.0E-06A   Avg: -500.1E-06   Err:    0.01%

%FS: -10.0   I: -100.0E-06A   Avg: -99.74E-06   Err:   -0.26%

%FS:  10.0   I:  100.0E-06A   Avg:  99.67E-06   Err:   -0.33%

%FS:  50.0   I:  500.0E-06A   Avg:  499.7E-06   Err:   -0.07%

%FS:  90.0   I:  900.0E-06A   Avg:  899.7E-06   Err:   -0.03%

Range  2   FullScale:    100.0E-06   Offset perc/FS:  0.0%

%FS: -90.0   I: -90.00E-06A   Avg: -90.05E-06   Err:    0.06%

%FS: -50.0   I: -50.00E-06A   Avg: -50.00E-06   Err:    0.01%

%FS: -10.0   I: -10.00E-06A   Avg: -9.955E-06   Err:   -0.45%

%FS:  10.0   I:  10.00E-06A   Avg:  10.02E-06   Err:    0.24%

%FS:  50.0   I:  50.00E-06A   Avg:  50.06E-06   Err:    0.12%

%FS:  90.0   I:  90.00E-06A   Avg:  90.11E-06   Err:    0.12%

Range  4   FullScale:    10.00E-06   Offset perc/FS:  0.0%

%FS: -90.0   I: -9.000E-06A   Avg: -9.004E-06   Err:    0.04%

%FS: -50.0   I: -5.000E-06A   Avg: -5.003E-06   Err:    0.06%

%FS: -10.0   I: -1.000E-06A   Avg: -999.1E-09   Err:   -0.09%

%FS:  10.0   I:  1.000E-06A   Avg:  997.7E-09   Err:   -0.23%

%FS:  50.0   I:  5.000E-06A   Avg:  4.997E-06   Err:   -0.06%

%FS:  90.0   I:  9.000E-06A   Avg:  9.000E-06   Err:   -0.00%

Range  6   FullScale:    1.000E-06   Offset perc/FS:  1.8%

%FS: -90.0   I: -900.0E-09A   Avg: -899.3E-09   Err:   -0.08%

%FS: -50.0   I: -500.0E-09A   Avg: -498.4E-09   Err:   -0.31%

%FS: -10.0   I: -100.0E-09A   Avg: -98.14E-09   Err:   -1.86%

%FS:  10.0   I:  100.0E-09A   Avg:  101.8E-09   Err:    1.79%

%FS:  50.0   I:  500.0E-09A   Avg:  501.9E-09   Err:    0.37%

%FS:  90.0   I:  900.0E-09A   Avg:  902.3E-09   Err:    0.25%

Range  8   FullScale:    100.0E-09   Offset perc/FS:  0.0%

%FS: -90.0   I: -90.00E-09A   Avg: -90.03E-09   Err:    0.04%

%FS: -50.0   I: -50.00E-09A   Avg: -49.99E-09   Err:   -0.03%

%FS: -10.0   I: -10.00E-09A   Avg: -9.971E-09   Err:   -0.29%

%FS:  10.0   I:  10.00E-09A   Avg:  10.01E-09   Err:    0.09%

%FS:  50.0   I:  50.00E-09A   Avg:  50.02E-09   Err:    0.05%

%FS:  90.0   I:  90.00E-09A   Avg:  90.03E-09   Err:    0.04%

Range 10   FullScale:    10.00E-09   Offset perc/FS:  0.1%

%FS: -90.0   I: -9.000E-09A   Avg: -9.006E-09   Err:    0.07%

%FS: -50.0   I: -5.000E-09A   Avg: -5.008E-09   Err:    0.15%

%FS: -10.0   I: -1.000E-09A   Avg: -1.007E-09   Err:    0.66%

%FS:  10.0   I:  1.000E-09A   Avg:  993.6E-12   Err:   -0.64%

%FS:  50.0   I:  5.000E-09A   Avg:  4.985E-09   Err:   -0.30%

%FS:  90.0   I:  9.000E-09A   Avg:  8.986E-09   Err:   -0.15%

Range 12   FullScale:    1.000E-09   Offset perc/FS:  2.6%

%FS: -90.0   I: -900.0E-12A   Avg: -900.0E-12   Err:   -0.00%

%FS: -50.0   I: -500.0E-12A   Avg: -500.3E-12   Err:    0.07%

%FS: -10.0   I: -100.0E-12A   Avg: -99.92E-12   Err:   -0.08%

%FS:  10.0   I:  100.0E-12A   Avg:  99.61E-12   Err:   -0.39%

%FS:  50.0   I:  500.0E-12A   Avg:  498.9E-12   Err:   -0.23%

%FS:  90.0   I:  900.0E-12A   Avg:  898.3E-12   Err:   -0.19%

Range 14   FullScale:    100.0E-12   Offset perc/FS: 30.6%

%FS: -90.0   I: -90.00E-12A   Avg: -90.34E-12   Err:    0.38%

%FS: -50.0   I: -50.00E-12A   Avg: -50.60E-12   Err:    1.20%

%FS: -10.0   I: -10.00E-12A   Avg: -10.68E-12   Err:    6.76%

%FS:  10.0   I:  10.00E-12A   Avg:  9.416E-12   Err:   -5.84%

%FS:  50.0   I:  50.00E-12A   Avg:  49.59E-12   Err:   -0.82%

%FS:  90.0   I:  90.00E-12A   Avg:  89.29E-12   Err:   -0.78%

Test stopped

Test time: 8 min 7 sec